Ultrasonic Flaw Detector

QUAM UT900

A、Introduction

Digital ultrasonic flaw detector is advanced type, which is touch screen, can quickly, easily and without damage, accurately detect, locate, evaluate and diagnose various defects inside the work-piece such as cracks, welds, pores, sand holes, inclusions, folding, etc. It has been applied for electric power, petrochemical, boiler and pressure vessel, steel structure, military, aerospace, railway transportation, automobile, machinery and other fields. It is an essential instrument for the non-destructive testing industry.

B、Features

●High-precision quantification and positioning to meet the requirements of near and far distance detection

●The near-field blind zone is small, can meet the detecting requirements for small-diameter and thin-walled pipe.

●Auto calibration function:one-button auto calibration,easy to operate, automatic test probe “P Delay”,”K value”,”X value” and the velocity

●Automatic display the defect echo position(Depth:d, Horizontal:p,Distance:s, Amplitude,dB,ф )

●Switch three scales freely(Depth:d, Horizontal:p,Distance:s)

●Auto gain, peak envelope, peak memory functions, which can improve the detection efficiency.

●Automatically record the flaw detection process and dynamic playback

●φ value calculation：Forging flaw detection by straight probe, can find the highest wave to conversion φ value automatically

●500 independent channels(can be expandable), which can input and store the detection standards of any industries freely, do not need to carry the standard blocks for on-site inspection.

●Store, playback 500 A-scan waves and data freely

● The DAC, AVG, and TCG curves (depth compensation) are automatically generated and can be segmented. The sampling points are unrestricted and can be corrected and compensated.

●14 built-in inspection standards

●Free to enter any industry standards

●Pulse width and strength can adjustable

●B scan and B color scan function；

●Can communicate with the computer, and export WORD. File, also the detection report

●IP65 ABS plastic case, sturdy and durable,water-proof and dust-proof, and excellent anti-interference ability

●Use (wireless) communication software to analyze data print reports, etc；

●260,000-color true color screen, is suitable for working environment under strong light or low light

●High performance lithium battery, can work continuously for 8-10 hours

●Real-time clock recording: real-time flaw detection date, time tracking record, and storage

●Power-down protection, storage data can not lost

●Flaw detection parameters can be automatically tested or preset

●Digital reject, does not affect gain and linearity

●Gain compensation: Db attenuation can be corrected for surface roughness, curved surfaces, long-range flaw detection of thick work-pieces, etc.

●Can operate the flaw detector by software at the PC, achieving the goal of computer-controlled flaw detector to detecting

C．Technical Specification

| Display | 7 inch TFT color screen，800*480 resolution |
| --- | --- |
| Operation mode | Button, Rotary, Touch Screen |
| Power supply | Lithium Ion Battery |
| Battery capacity | 5.0Ah |
| Power voltage | 12V |
| Power quantity | 1 |
| Working time | ≥8 hours |
| Adapter input | DC100～240V 50Hz/60Hz |
| Adapter output | AC 12V |
| Adapter power | 36W |
| Data storage | SD card(16G) |
| Alarm | 1 |
| Working Temperature | -10℃～45℃ |
| Storage Temperature | -20℃～60℃ |
| IP Grade | IP65 |
| Dimension | 245*155*55mm |
| Weight | 1.18kg（included battery） |
| Conventional UT Model |  |
| Probe Connector type | LEMO 00 |
| Channel Type | Single channel |
| Channel Num | 500 group(able to be scaled) |
| Pulse Type | Negative sharp wave |
| Transmit Voltage | 50～350V，step in 50V |
| Damping | 560Ω |
| Gain | 0~110dB，step：0.5/2/6/12dB |
| Gain Fine Adjustment | -4~+4 |
| Surface compensation | All Gain Range |
| Working Frequency | 0.5~20MHz； |
| Probe Type | Single，Dual, Through, Immersion type |
| Filter | Three optional: 1~4MHz/0.5~10MHz/2~20MHz |
| Detection Mode | Negative/Positive/Two-way/RF |
| Reject | 0~80%，step 1% |
| Testing Range | 0~15000mm，Minimum display range:15mm（in steel） |
| Material Velocity | 100~20000m/s |
| Pulse displacement | -10~1000mm |
| P DELAY | 0~200us |
| X-VAL | 0~100mm |
| Guide | Weld, Sheet, Forging inspection |
| Testing Point | Peak/X-val/J val |
| Measurement | Gate：Amplitude、Amplitude dB value、Range、Horizontal distance、Vertical distance、The difference value between A and B Gate Cursor:2 cross cursors, can test the horizontal and vertical distance, and the distance between cursors(under B scan function) |
| Gate | Gate start:all range |
|  | Gate width:all range |
|  | Gate level：10~90%，step:1% |
| Curve | DAC, maximum six curves, meet to NB/T 47013,GB/T 11345,GB/T 29712, and other standards |
|  | TCG, maximum six curves |
|  | AVG |
| Other functions | Full screen, cursor switch(range/Height/Horizon), single/continuous auto gain(10~100%, step 10%), echo compare,echo full, peak envelope,peak memory,fast scan,outside mode,screenshot |
|  | Peak freeze/Crack depth/Gate expansion//curved surface modification/ B scan/Flat weld simulation//video |
| Alarm | Sound and light alarm |
| Sensitivity Leavings | ≥65dB(200mm—Φ2FH,2.5PΦ20) |
| Horizontal linearity error | ≤0.3% |
| Vertical linearity error | ≤3% |
| Amplitude linearity error | ≤±2% |
| Attenuator accuracy | 20dB ±1 dB |
| Dynamic range | ≥32dB |
| Distant resolution | ≥26dB |
| Noise Level | ＜40×10-9 V/ |

D、Standard Configuration

|  | Name | QTY |
| --- | --- | --- |
| 1 | Main unit | 1 |
| 2 | Power adaptor | 1 |
| 3 | probe connecting cable | 2 |
| 4 | Instrument case | 1 |
| 5 | Instruction manual | 1 |
| 6 | Straight probe 10mm 2.5MHz | 1 |
| 7 | Angle probe 13x13 K2 2.5MHz | 1 |
| 8 | Warranty card | 1 |

ASIA TRINITY ENTERPRISES LIMITED

Address: Rm.1405.1407, Tianle Building, 1021 Buji Road, LoHu, Shenzhen.

Tel： (86) 0755-2394-2087

Fox：(86) 0755-2390-7646

Mail：info@asiatrinity.com

Web：www.quamtec.com